## SOLUTIONS_TEXTE
Mot_Cle | Type | Priorite | Cause | Solution
4A01 | Thermal | MOYENNE | Système de refroidissement insuffisant ou filtre encrassé | 1. Vérifier ventilation salle 2. Nettoyer filtres 3. Contrôler liquide de refroidissement
4B02 | Hardware | HAUTE | Usure roulement gantry ou déséquilibre mécanique | 1. Arrêter acquisitions 2. Inspecter roulements 3. Recalibrer vitesse rotation
7F03 | Power | HAUTE | Défaillance module alimentation haute tension | 1. Arrêt immédiat 2. Vérifier fusibles HT 3. Tester condensateurs 4. Appeler SAV constructeur
3C04 | Calibration | MOYENNE | Dérive des paramètres détecteur après usage prolongé | 1. Lancer séquence calibration constructeur 2. Vérifier fantôme QC 3. Documenter résultats
5A01 | Hardware | MOYENNE | Fixation bobine gradient desserrée ou amortisseur usé | 1. Vérifier serrage fixations 2. Inspecter amortisseurs 3. Planifier maintenance si récurrent
5B02 | Hardware | HAUTE | Connecteur RF dégradé ou interférence électromagnétique | 1. Inspecter connecteurs RF 2. Vérifier blindage cage Faraday 3. Recalibrer chaîne RF
6C03 | Network | MOYENNE | Câble réseau défectueux ou surcharge réseau DICOM | 1. Tester câble Ethernet 2. Vérifier switch réseau 3. Redémarrer service DICOM
8D01 | Tube RX | HAUTE | Tube radiogène en fin de vie (heures d'utilisation élevées) | 1. Vérifier compteur heures tube 2. Commander tube de remplacement 3. Planifier changement
0301H | TubeSpits | 2 | ?
0302H | TubeSpits | 2 | ?
0303H | TubeSpits | 2 | ?
0304H | TubeSpits | 2 | ?
0305H | TubeSpits | 2 | ?
0324H | TubeSpits | 2 | ?
0328H | TubeSpits | 1 | ? | No action
0101H | Rotation | 4 | ?
0102H | Rotation | 4 | ?
0103H | Rotation | 4 | ?
0104H | Rotation | 4 | ?
0105H | Rotation | 4 | ?
0106H | Rotation | 4 | ?
0107H | Rotation | 4 | ?
0108H | Rotation | 4 | ?
0109H | Rotation | 4 | ?
0110H | Rotation | 4 | ?
0111H | Rotation | 4 | ?
0112H | Rotation | 4 | ?
0113H | Rotation | 4 | ?
0114H | Rotation | 4 | ?
0115H | Rotation | 4 | ?
0149H | Rotation | 4 | ?
0201H | Heater0 | 4 | ?
0202H | Heater0 | 4 | ?
0203H | Heater0 | 4 | ?
0204H | Heater0 | 4 | ?
0205H | Heater0 | 4 | ?
0206H | Heater0 | 4 | ?
0207H | Heater0 | 4 | ?
0208H | Heater0 | 4 | ?
0209H | Heater0 | 4 | ?
0210H | Heater0 | 4 | ?
0211H | Heater0 | 4 | ?
0212H | Heater0 | 4 | ?
0213H | Heater0 | 4 | ?
0214H | Heater0 | 4 | ?
0215H | Heater0 | 4 | ?
0216H | Heater0 | 4 | ?
0217H | Heater0 | 4 | ?
0221H | Heater0 | 4 | ?
0222H | Heater0 | 4 | ?
0223H | Heater0 | 4 | ?
0224H | Heater0 | 4 | ?
0248H | Heater0 | 4 | ?
0249H | Heater0 | 4 | ?
1201H | Heater1 | 4 | ?
1202H | Heater1 | 4 | ?
1203H | Heater1 | 4 | ?
1204H | Heater1 | 4 | ?
1205H | Heater1 | 4 | ?
1206H | Heater1 | 4 | ?
1207H | Heater1 | 4 | ?
1208H | Heater1 | 4 | ?
1209H | Heater1 | 4 | ?
1210H | Heater1 | 4 | ?
1211H | Heater1 | 4 | ?
1212H | Heater1 | 4 | ?
1213H | Heater1 | 4 | ?
1214H | Heater1 | 4 | ?
1215H | Heater1 | 4 | ?
1216H | Heater1 | 4 | ?
1217H | Heater1 | 4 | ?
1221H | Heater1 | 4 | ?
1222H | Heater1 | 4 | ?
1223H | Heater1 | 4 | ?
1224H | Heater1 | 4 | ?
0306H | Exposure | 3 | ?
0307H | Exposure | 3 | ?
0308H | Exposure | 3 | ?
0309H | Exposure | 3 | ?
0310H | Exposure | 3 | ?
0311H | Exposure | 3 | ?
0312H | Exposure | 3 | ?
0313H | Exposure | 3 | ?
0314H | Exposure | 3 | ?
0316H | Exposure | 3 | ?
0317H | Exposure | 3 | ?
0318H | Exposure | 3 | ?
0319H | Exposure | 3 | ?
0320H | Exposure | 3 | ?
0323H | Exposure | 3 | ?
0325H | Exposure | 3 | ?
0326H | Exposure | 3 | ?
0327H | Exposure | 3 | ?
0329H | Exposure | 3 | ?
0501H | PowerSupply | 4 | ?
0502H | PowerSupply | 4 | ?
0503H | PowerSupply | 4 | ?
0505H | PowerSupply | 4 | ?
0506H | PowerSupply | 4 | ?
0507H | PowerSupply | 4 | ?
0510H | PowerSupply | 3 | ?
0511H | PowerSupply | 3 | ?
0549H | PowerSupply | 4 | ?
0553H | PowerSupply | 4 | ?
0557H | PowerSupply | 4 | ?
0563H | PowerSupply | 4 | ?
0567H | PowerSupply | 4 | ?
0573H | PowerSupply | 4 | ?
0577H | PowerSupply | 4 | ?
0180H | Hardware | 4 | ?
0181H | Hardware | 4 | ?
0280H | Hardware | 4 | ?
0281H | Hardware | 4 | ?
0322H | Hardware | 4 | ?
03A0H | Hardware | 4 | ?
0601H | Hardware | 4 | ?
0602H | Hardware | 4 | ?
0902H | Hardware | 4 | ?
1000H | Hardware | 4 | ?
1006H | Hardware | 4 | ?
1007H | Hardware | 4 | ?
1008H | Hardware | 4 | ?
1010H | Hardware | 4 | ?
1180H | Hardware | 4 | ?
1182H | Hardware | 4 | ?
1280H | Hardware | 4 | ?
1281H | Hardware | 4 | ?
1402H | Hardware | 4 | ?
1403H | Hardware | 4 | ?
1404H | Hardware | 4 | ?
1413H | Hardware | 4 | ?
1414H | Hardware | 4 | ?
1415H | Hardware | 4 | ?
0701H | Application | 4 | ?
0702H | Application | 4 | ?
0705H | Application | 4 | ?
0706H | Application | 4 | ?
0707H | Application | 4 | ?
1012H | Application | 4 | ?
1424H | Application | 4 | ?
0603H | Communication | 1 | ? | No action
0605H | Communication | 1 | ? | No action
0606H | Communication | 1 | ? | No action
1301H | Communication | 4 | ?
1302H | Communication | 4 | ?
0804H | Thermal | 5 | ?
0805H | Thermal | 5 | ?
0806H | Thermal | 5 | ?
0807H | Thermal | 1 | ? | No action
0808H | Thermal | 1 | ? | No action
0903H | Thermal | 5 | ?
0905H | Thermal | 5 | ?
1405H | Thermal | 5 | ?
1406H | Thermal | 5 | ?
1412H | Thermal | 5 | ?
1500H | Manipulation | 3 | ?
1501H | Manipulation | 3 | ?
1502H | Manipulation | 3 | ?
0151H | RotationWarning | 1 | ? | No action
0152H | RotationWarning | 1 | ? | No action
0153H | RotationWarning | 1 | ? | No action
0154H | RotationWarning | 1 | ? | No action
0155H | RotationWarning | 1 | ? | No action
0156H | RotationWarning | 1 | ? | No action
0157H | RotationWarning | 1 | ? | No action
0158H | RotationWarning | 1 | ? | No action
0159H | RotationWarning | 1 | ? | No action
0160H | RotationWarning | 1 | ? | No action
0161H | RotationWarning | 1 | ? | No action
0162H | RotationWarning | 1 | ? | No action
0163H | RotationWarning | 1 | ? | No action
0164H | RotationWarning | 1 | ? | No action
0199H | RotationWarning | 1 | ? | No action
0330H | RotationWarning | 1 | ? | No action
0251H | HeaterWarning | 1 | ? | No action
0252H | HeaterWarning | 1 | ? | No action
0253H | HeaterWarning | 1 | ? | No action
0254H | HeaterWarning | 1 | ? | No action
0255H | HeaterWarning | 1 | ? | No action
0256H | HeaterWarning | 1 | ? | No action
0257H | HeaterWarning | 1 | ? | No action
0258H | HeaterWarning | 1 | ? | No action
0259H | HeaterWarning | 1 | ? | No action
0260H | HeaterWarning | 1 | ? | No action
0261H | HeaterWarning | 1 | ? | No action
0262H | HeaterWarning | 1 | ? | No action
0263H | HeaterWarning | 1 | ? | No action
0299H | HeaterWarning | 1 | ? | No action
1251H | HeaterWarning | 1 | ? | No action
1252H | HeaterWarning | 1 | ? | No action
1253H | HeaterWarning | 1 | ? | No action
1254H | HeaterWarning | 1 | ? | No action
1255H | HeaterWarning | 1 | ? | No action
1256H | HeaterWarning | 1 | ? | No action
1257H | HeaterWarning | 1 | ? | No action
1258H | HeaterWarning | 1 | ? | No action
1259H | HeaterWarning | 1 | ? | No action
1260H | HeaterWarning | 1 | ? | No action
1261H | HeaterWarning | 1 | ? | No action
1262H | HeaterWarning | 1 | ? | No action
1263H | HeaterWarning | 1 | ? | No action
1299H | HeaterWarning | 1 | ? | No action
0508H | LVPSWarning | 1 | ? | No action
0509H | LVPSWarning | 1 | ? | No action
0404H | ApplicationWarning | 2 | ? | No action
0703H | ApplicationWarning | 1 | ? | No action
0704H | ApplicationWarning | 2 | ? | No action
0708H | ApplicationWarning | 1 | ? | No action
0810H | ApplicationWarning | 1 | ? | No action
0811H | ApplicationWarning | 1 | ? | No action
0907H | ApplicationWarning | 1 | ? | No action
1401H | ApplicationWarning | 1 | ? | Replace battery of PPC KVCTRL board

## CODES_HEXA
Code | Message | Niveau | Type
4A01 | Température tube élevée | ATTENTION | Thermal
4B02 | Erreur rotation gantry | ERREUR | Hardware
7F03 | Défaillance alimentation HT | CRITIQUE | Power
3C04 | Calibration détecteur requise | ATTENTION | Calibration
5A01 | Vibration bobine gradient | ATTENTION | Hardware
5B02 | Perte signal RF intermittente | ERREUR | Hardware
6C03 | Erreur communication console | ERREUR | Network
8D01 | Usure tube RX détectée | ATTENTION | Tube RX
0301H | Tube spit (kV+ and kV- dropped) | ATTENTION | TubeSpits
0302H | Tube spit (kV+ has dropped) | ATTENTION | TubeSpits
0303H | Tube spit (kV- has dropped) | ATTENTION | TubeSpits
0304H | Tube spit (kV regulation error) | ATTENTION | TubeSpits
0305H | FPGA problem (restarting safety signal) | ATTENTION | TubeSpits
0324H | Tube spit (all kV drop/regul errors) | ATTENTION | TubeSpits
0328H | Spit counter overflow | ATTENTION | TubeSpits
0101H | No CAN message received within 5 sec | ATTENTION | Rotation
0102H | Database not correct | ATTENTION | Rotation
0103H | Current overload | ATTENTION | Rotation
0104H | Open load | ATTENTION | Rotation
0105H | Phases unbalanced | ATTENTION | Rotation
0106H | Phases error | ATTENTION | Rotation
0107H | Inverter permanent overcurrent | ATTENTION | Rotation
0108H | Command error | ATTENTION | Rotation
0109H | MAINS_DROP has failed | ATTENTION | Rotation
0110H | PRD error (Z3/Z4=bitmap) | ATTENTION | Rotation
0111H | F0 main frequency problem | ATTENTION | Rotation
0112H | HW/FW config error | ATTENTION | Rotation
0113H | IUVW short circuit error | ATTENTION | Rotation
0114H | Rotor not at speed | ATTENTION | Rotation
0115H | HV cable short circuit error | ATTENTION | Rotation
0149H | Unknown rotation error | ATTENTION | Rotation
0201H | Heater 0: no CAN message received within 5 secs | ATTENTION | Heater0
0202H | Heater 0: database not correct | ATTENTION | Heater0
0203H | Heater 0: inverter permanent overcurrent | ATTENTION | Heater0
0204H | Heater 0: filament permanent open circuit | ATTENTION | Heater0
0205H | Heater 0: inverter permanent short circuit | ATTENTION | Heater0
0206H | Heater 0: filament overload Preâ€‘Heat | ATTENTION | Heater0
0207H | Heater 0: filament overload Boost | ATTENTION | Heater0
0208H | Heater 0: filament overload Heat | ATTENTION | Heater0
0209H | Heater 0: command error | ATTENTION | Heater0
0210H | Heater 0: PWM rate too low | ATTENTION | Heater0
0211H | Heater 0: PWM rate too high | ATTENTION | Heater0
0212H | Heater 0: MAINS_DROP detected | ATTENTION | Heater0
0213H | Heater 0: PRD error | ATTENTION | Heater0
0214H | Heater 0: boost too long | ATTENTION | Heater0
0215H | Heater 0: filament selection error | ATTENTION | Heater0
0216H | Heater 0: current feedback not null when inverter OFF | ATTENTION | Heater0
0217H | Heater 0: current not zero through a not selected filament | ATTENTION | Heater0
0221H | Heater 0: filament Database not correct (tube1) | ATTENTION | Heater0
0222H | Heater 0: filament Database not correct (tube2) | ATTENTION | Heater0
0223H | Heater 0: filament Database not correct (tube3) | ATTENTION | Heater0
0224H | Heater 0: filament Database not correct (tube4) | ATTENTION | Heater0
0248H | Heater large focus error | ATTENTION | Heater0
0249H | Heater small focus error | ATTENTION | Heater0
1201H | Heater 1: no CAN message received within 5 secs | ATTENTION | Heater1
1202H | Heater 1: database not correct | ATTENTION | Heater1
1203H | Heater 1: inverter permanent overcurrent | ATTENTION | Heater1
1204H | Heater 1: filament permanent open circuit | ATTENTION | Heater1
1205H | Heater 1: inverter permanent short circuit | ATTENTION | Heater1
1206H | Heater 1: filament overload Preâ€‘Heat | ATTENTION | Heater1
1207H | Heater 1: filament overload Boost | ATTENTION | Heater1
1208H | Heater 1: filament overload Heat | ATTENTION | Heater1
1209H | Heater 1: command error | ATTENTION | Heater1
1210H | Heater 1: PWM rate too low | ATTENTION | Heater1
1211H | Heater 1: PWM rate too high | ATTENTION | Heater1
1212H | Heater 1: MAINS_DROP detected | ATTENTION | Heater1
1213H | Heater 1: PRD error | ATTENTION | Heater1
1214H | Heater 1: boost too long | ATTENTION | Heater1
1215H | Heater 1: filament selection error | ATTENTION | Heater1
1216H | Heater 1: current feedback not null when inverter OFF | ATTENTION | Heater1
1217H | Heater 1: current not zero through a not selected filament | ATTENTION | Heater1
1221H | Heater 1: filament Database not correct (tube1) | ATTENTION | Heater1
1222H | Heater 1: filament Database not correct (tube2) | ATTENTION | Heater1
1223H | Heater 1: filament Database not correct (tube3) | ATTENTION | Heater1
1224H | Heater 1: filament Database not correct (tube4) | ATTENTION | Heater1
0306H | No kV feedback on anode side | ATTENTION | Exposure
0307H | No kV feedback on cathode side | ATTENTION | Exposure
0308H | No kV feedback on anode and cathode | ATTENTION | Exposure
0309H | kV detected during kV diagnostics | ATTENTION | Exposure
0310H | kV max detected | ATTENTION | Exposure
0311H | ILP current not OK | ATTENTION | Exposure
0312H | ILR current not OK | ATTENTION | Exposure
0313H | Inverter max ILR current detected | ATTENTION | Exposure
0314H | ILR current timeout | ATTENTION | Exposure
0316H | Spit Max error | ATTENTION | Exposure
0317H | Spit Ratio error | ATTENTION | Exposure
0318H | kV did not reach 75% after 20â€¯ms | ATTENTION | Exposure
0319H | kV unbalanced detected | ATTENTION | Exposure
0320H | FPGA problem; safety hit with unknown reason | ATTENTION | Exposure
0323H | ILP and ILR current not OK | ATTENTION | Exposure
0325H | DSP overload | ATTENTION | Exposure
0326H | DSP exposure stop | ATTENTION | Exposure
0327H | DSP regulation out of range | ATTENTION | Exposure
0329H | Wrong HV capa measure | ATTENTION | Exposure
0501H | DC bus out of range (<400â€¯Vâ€¯orâ€¯>â€¯850â€¯V) | ATTENTION | PowerSupply
0502H | Error Overvoltageâ€¯>â€¯780â€¯VDC | ATTENTION | PowerSupply
0503H | Inverter Gate Power Supply error (checked at prep) | ATTENTION | PowerSupply
0505H | Mains power supply has dropped during exposure | ATTENTION | PowerSupply
0506H | DC busâ€¯1â€‘phase preâ€‘charge error | ATTENTION | PowerSupply
0507H | DC busâ€¯1â€‘phase discharge error | ATTENTION | PowerSupply
0510H | HVPM AEC notâ€¯230â€¯V | ATTENTION | PowerSupply
0511H | Undervoltageâ€¯<â€¯320â€¯VDC | ATTENTION | PowerSupply
0549H | LVPSâ€¯TRI: Unknown error | ATTENTION | PowerSupply
0553H | LVPSâ€¯TRI: Detectedâ€¯+160â€¯V too high | ATTENTION | PowerSupply
0557H | LVPSâ€¯TRI: Detectedâ€¯+160â€¯V too low | ATTENTION | PowerSupply
0563H | LVPSâ€¯TRI: Detectedâ€¯+15â€¯V too high | ATTENTION | PowerSupply
0567H | LVPSâ€¯TRI: Detectedâ€¯+15â€¯V too low | ATTENTION | PowerSupply
0573H | LVPSâ€¯TRI: Detectedâ€¯â€“15â€¯V too high | ATTENTION | PowerSupply
0577H | LVPSâ€¯TRI: Detectedâ€¯â€“15â€¯V too low | ATTENTION | PowerSupply
0180H | Rotation board communication problem | ATTENTION | Hardware
0181H | Rotation board has reset | ATTENTION | Hardware
0280H | Heater board communication problem | ATTENTION | Hardware
0281H | Heater board has reset | ATTENTION | Hardware
0322H | kV ref ADC / DAC failed | ATTENTION | Hardware
03A0H | DSP error | ATTENTION | Hardware
0601H | RTL error (+ associated data to check which of the 4 RTL lines) | ATTENTION | Hardware
0602H | External CAN bus off | ATTENTION | Hardware
0902H | Tube cooling supply error | ATTENTION | Hardware
1000H | Tube Switch communication problem | ATTENTION | Hardware
1006H | HV switch selection error | ATTENTION | Hardware
1007H | Rotor switch selection error | ATTENTION | Hardware
1008H | Heater switch selection error | ATTENTION | Hardware
1010H | Tube Switch board has reset | ATTENTION | Hardware
1180H | Ingrid not active | ATTENTION | Hardware
1182H | Ingrid CAN bus off | ATTENTION | Hardware
1280H | Heaterâ€¯1 board communication problem | ATTENTION | Hardware
1281H | Heaterâ€¯1 board has reset | ATTENTION | Hardware
1402H | Internal CAN bus off | ATTENTION | Hardware
1403H | Connection fault | ATTENTION | Hardware
1404H | FPGA configuration problem | ATTENTION | Hardware
1413H | Internal CAN bus Txo | ATTENTION | Hardware
1414H | Internal CAN LD bus Txo | ATTENTION | Hardware
1415H | External CAN bus Txo | ATTENTION | Hardware
0701H | NV RAM checksum problem | ATTENTION | Application
0702H | Software problem | ATTENTION | Application
0705H | System or database configuration error | ATTENTION | Application
0706H | Wrong DSP version | ATTENTION | Application
0707H | Wrong FPGA version | ATTENTION | Application
1012H | Exposure preparation while tube not ready | ATTENTION | Application
1424H | Thermal algorithm inconsistent with database | ATTENTION | Application
0603H | Debug screen communication error | ATTENTION | Communication
0605H | TAV communication error | ATTENTION | Communication
0606H | MPC/Madrid communication problem | ATTENTION | Communication
1301H | AEC Board communication problem | ATTENTION | Communication
1302H | Watchdog CAN communication error | ATTENTION | Communication
0804H | Tank thermal sensor too high | ATTENTION | Thermal
0805H | Inverter HW temperature too high | ATTENTION | Thermal
0806H | Lp temperature too high | ATTENTION | Thermal
0807H | Tank thermal sensor warning | ATTENTION | Thermal
0808H | Inv thermal sensor warning | ATTENTION | Thermal
0903H | Tube exceeded 70â€¯Â°C | ATTENTION | Thermal
0905H | Tube thermal protection | ATTENTION | Thermal
1405H | Tank temperature sensor too low | ATTENTION | Thermal
1406H | Inverter HW temperature measurement too low | ATTENTION | Thermal
1412H | Lp temperature too low | ATTENTION | Thermal
1500H | Tomo brightness not good | ATTENTION | Manipulation
1501H | Release exposure switch (manipulator error) | ATTENTION | Manipulation
1502H | AEC does not cut exposure | ATTENTION | Manipulation
0151H | CAN Domain command number error | ATTENTION | RotationWarning
0152H | CAN Domain request with no transfer init | ATTENTION | RotationWarning
0153H | CAN Domain Toggle bit error | ATTENTION | RotationWarning
0154H | CAN Domain : less than 2 data to download | ATTENTION | RotationWarning
0155H | CAN Domain Abort received & applied | ATTENTION | RotationWarning
0156H | Bad index in config upload | ATTENTION | RotationWarning
0157H | Tube switch while rotor not off | ATTENTION | RotationWarning
0158H | Acceleration cmd while no tube selected | ATTENTION | RotationWarning
0159H | Acceleration cmd while database not OK | ATTENTION | RotationWarning
0160H | Database download while rotor speeding | ATTENTION | RotationWarning
0161H | Acceleration command not OK | ATTENTION | RotationWarning
0162H | Acceleration while in error | ATTENTION | RotationWarning
0163H | No CAN message received within 4 sec | ATTENTION | RotationWarning
0164H | Inverter overcurrent (<â€¯3â€¯times) | ATTENTION | RotationWarning
0199H | Unknown warning | ATTENTION | RotationWarning
0330H | Chiller warning | ATTENTION | RotationWarning
0251H | Heater 0: received command is not OK | ATTENTION | HeaterWarning
0252H | Heater 0: command not OK | ATTENTION | HeaterWarning
0253H | Heater 0: no CAN message received within 4 secs | ATTENTION | HeaterWarning
0254H | Heater 0: inverter overcurrent (inverter1) (<â€¯3â€¯times) | ATTENTION | HeaterWarning
0255H | Heater 0: filament open circuit (inverter1) (<â€¯3â€¯times) | ATTENTION | HeaterWarning
0256H | Heater 0: inverter short circuit (inverter1) (<â€¯3â€¯times) | ATTENTION | HeaterWarning
0257H | Heater 0: tube switch while filaments not OFF | ATTENTION | HeaterWarning
0258H | Heater 0: CAN Domain command number error | ATTENTION | HeaterWarning
0259H | Heater 0: CAN Domain request with no transfer init | ATTENTION | HeaterWarning
0260H | Heater 0: CAN Domain Toggle bit error | ATTENTION | HeaterWarning
0261H | Heater 0: CAN Domain less than 2 data to download | ATTENTION | HeaterWarning
0262H | Heater 0: CAN Domain Abort received & applied | ATTENTION | HeaterWarning
0263H | Heater 0: database download while heater not cut | ATTENTION | HeaterWarning
0299H | Heater 0: unknown warning | ATTENTION | HeaterWarning
1251H | Heater 1: received command is not OK | ATTENTION | HeaterWarning
1252H | Heater 1: command not OK | ATTENTION | HeaterWarning
1253H | Heater 1: no CAN message received within 4 secs | ATTENTION | HeaterWarning
1254H | Heater 1: inverter overcurrent (inverter1) (<â€¯3â€¯times) | ATTENTION | HeaterWarning
1255H | Heater 1: filament open circuit (inverter1) (<â€¯3â€¯times) | ATTENTION | HeaterWarning
1256H | Heater 1: inverter short circuit (inverter1) (<â€¯3â€¯times) | ATTENTION | HeaterWarning
1257H | Heater 1: tube switch while filaments not OFF | ATTENTION | HeaterWarning
1258H | Heater 1: CAN Domain command number error | ATTENTION | HeaterWarning
1259H | Heater 1: CAN Domain request with no transfer init | ATTENTION | HeaterWarning
1260H | Heater 1: CAN Domain Toggle bit error | ATTENTION | HeaterWarning
1261H | Heater 1: CAN Domain : less than 2 data to download | ATTENTION | HeaterWarning
1262H | Heater 1: CAN Domain Abort received & applied | ATTENTION | HeaterWarning
1263H | Heater 1: database download while heater not cut | ATTENTION | HeaterWarning
1299H | Heater 1: unknown warning | ATTENTION | HeaterWarning
0508H | Warning Overvoltageâ€¯>â€¯755â€¯VDC | ATTENTION | LVPSWarning
0509H | Warning Undervoltageâ€¯<â€¯320â€¯VDC | ATTENTION | LVPSWarning
0404H | No Pulse Warning | ATTENTION | ApplicationWarning
0703H | NMI occured (MSB @) | ATTENTION | ApplicationWarning
0704H | Heater / Rotor hold too long | ATTENTION | ApplicationWarning
0708H | NMI occured (LSB @) | ATTENTION | ApplicationWarning
0810H | Configuration warning | ATTENTION | ApplicationWarning
0811H | Thermal switch warning | ATTENTION | ApplicationWarning
0907H | Filament aging saturated | ATTENTION | ApplicationWarning
1401H | Saved RAM battery life time limit reached | ATTENTION | ApplicationWarning
2026 | Erreur Inconnue | ATTENTION | Software
DDEFA | Erreur Inconnue | ATTENTION | Software
TXT | E:\Ana\JPN\HmcData\DcmService\Error_Jpn.txt | ATTENTION | Software

## HISTORIQUE
Date | Machine | Code | Type | Severite | Resolu
2025-11-17 14:02:58 | Radio_DR_02.log | 7F03 | Tube RX | BASSE | Non
2025-12-23 09:22:58 | Mammo_01.log | 5A01 | Tube RX | MOYENNE | Non
2025-12-15 21:15:58 | Scanner_CT_01.log | 7F03 | Détecteur | CRITIQUE | Non
2026-01-08 19:39:58 | Mammo_01.log | 6C03 | Calibration | HAUTE | Oui
2026-01-24 15:41:58 | Radio_DR_02.log | 4B02 | Hardware | BASSE | Non
2026-02-09 14:27:58 | Scanner_CT_01.log | 5B02 | Power | BASSE | Oui
2026-02-07 19:48:58 | Echo_Salle3.log | 6C03 | Network | BASSE | Oui
2026-01-31 04:12:58 | Scanner_CT_01.log | 5A01 | Hardware | CRITIQUE | Non
2025-11-12 13:51:58 | Scanner_CT_01.log | 4B02 | Détecteur | CRITIQUE | Oui
2025-12-28 18:35:58 | Mammo_01.log | 4B02 | Détecteur | CRITIQUE | Non
2025-11-22 18:03:58 | Mammo_01.log | 7F03 | Calibration | HAUTE | Non
2025-12-26 01:57:58 | Mammo_01.log | 6C03 | Tube RX | MOYENNE | Oui
2025-12-10 06:44:58 | Echo_Salle3.log | 7F03 | Network | CRITIQUE | Oui
2025-11-12 20:26:58 | Scanner_CT_01.log | 6C03 | Tube RX | BASSE | Non
2026-01-28 09:06:58 | IRM_Salle1.log | 5B02 | Tube RX | BASSE | Oui
2025-12-06 00:56:58 | Scanner_CT_01.log | 4B02 | Network | MOYENNE | Oui
2025-11-22 03:28:58 | Echo_Salle3.log | 3C04 | Hardware | CRITIQUE | Non
2026-01-31 16:42:58 | Scanner_CT_01.log | 6C03 | Hardware | MOYENNE | Non
2026-01-29 13:13:58 | Radio_DR_02.log | 8D01 | Détecteur | MOYENNE | Non
2025-12-20 07:51:58 | IRM_Salle1.log | 5A01 | Calibration | BASSE | Oui
2026-01-21 06:47:58 | Scanner_CT_01.log | 5A01 | Détecteur | MOYENNE | Non
2025-12-30 22:03:58 | IRM_Salle1.log | 5A01 | Hardware | BASSE | Non
2026-02-09 02:00:58 | Radio_DR_02.log | 7F03 | Tube RX | MOYENNE | Oui
2025-12-08 14:50:58 | Echo_Salle3.log | 7F03 | Calibration | HAUTE | Non
2026-01-11 14:43:58 | IRM_Salle1.log | 5A01 | Tube RX | CRITIQUE | Non
2025-11-21 05:46:58 | Mammo_01.log | 8D01 | Tube RX | MOYENNE | Oui
2026-01-12 08:10:58 | IRM_Salle1.log | 8D01 | Power | CRITIQUE | Non
2025-11-12 05:44:58 | Echo_Salle3.log | 3C04 | Détecteur | MOYENNE | Oui
2026-01-08 00:23:58 | Mammo_01.log | 8D01 | Network | HAUTE | Non
2026-02-09 03:24:58 | Echo_Salle3.log | 6C03 | Tube RX | CRITIQUE | Oui
2026-01-24 02:07:58 | Echo_Salle3.log | 8D01 | Hardware | CRITIQUE | Non
2025-12-13 06:14:58 | Radio_DR_02.log | 8D01 | Calibration | HAUTE | Non
2026-01-24 01:28:58 | Radio_DR_02.log | 6C03 | Network | HAUTE | Non
2025-11-16 00:45:58 | IRM_Salle1.log | 5B02 | Network | MOYENNE | Oui
2025-12-19 07:52:58 | Radio_DR_02.log | 8D01 | Software | HAUTE | Non
2026-01-09 05:56:58 | Scanner_CT_01.log | 5A01 | Détecteur | BASSE | Non
2026-02-05 05:52:58 | Scanner_CT_01.log | 5B02 | Détecteur | CRITIQUE | Non
2025-12-07 12:49:58 | IRM_Salle1.log | 3C04 | Hardware | HAUTE | Non
2025-11-12 16:11:58 | Echo_Salle3.log | 6C03 | Power | HAUTE | Oui
2025-11-24 00:47:58 | Mammo_01.log | 8D01 | Calibration | HAUTE | Non
2026-02-01 19:54:58 | IRM_Salle1.log | 4B02 | Software | BASSE | Non
2026-01-20 05:17:58 | Mammo_01.log | 5B02 | Software | BASSE | Non
2025-11-20 01:24:58 | IRM_Salle1.log | 5A01 | Software | BASSE | Non
2025-12-29 01:36:58 | Radio_DR_02.log | 6C03 | Software | MOYENNE | Oui
2026-01-03 03:37:58 | Scanner_CT_01.log | 4B02 | Tube RX | CRITIQUE | Oui
2026-01-15 04:41:58 | Radio_DR_02.log | 4B02 | Power | HAUTE | Oui
2025-12-22 10:15:58 | Scanner_CT_01.log | 6C03 | Hardware | CRITIQUE | Oui
2025-12-07 15:04:58 | IRM_Salle1.log | 4B02 | Network | MOYENNE | Oui
2025-11-19 16:54:58 | Scanner_CT_01.log | 4A01 | Hardware | BASSE | Non
2026-01-05 07:01:58 | Echo_Salle3.log | 3C04 | Network | CRITIQUE | Oui
2025-11-20 09:04:58 | Scanner_CT_01.log | 8D01 | Tube RX | BASSE | Oui
2025-11-20 18:29:58 | IRM_Salle1.log | 3C04 | Software | CRITIQUE | Oui
2025-12-23 03:19:58 | IRM_Salle1.log | 4B02 | Calibration | HAUTE | Oui
2026-02-09 13:18:58 | IRM_Salle1.log | 4A01 | Détecteur | CRITIQUE | Non
2026-01-25 09:45:58 | IRM_Salle1.log | 7F03 | Hardware | HAUTE | Non
2026-01-08 04:43:58 | IRM_Salle1.log | 7F03 | Détecteur | HAUTE | Oui
2026-02-09 23:19:58 | Scanner_CT_01.log | 8D01 | Software | MOYENNE | Oui
2025-12-09 03:26:58 | Radio_DR_02.log | 6C03 | Hardware | MOYENNE | Oui
2025-11-23 19:14:58 | Echo_Salle3.log | 6C03 | Software | HAUTE | Oui
2026-02-06 04:07:58 | Echo_Salle3.log | 6C03 | Power | MOYENNE | Oui
2026-01-14 05:54:58 | Mammo_01.log | 5A01 | Calibration | BASSE | Non
2025-12-22 16:15:58 | Scanner_CT_01.log | 5B02 | Calibration | HAUTE | Oui
2026-01-04 22:30:58 | IRM_Salle1.log | 5B02 | Network | BASSE | Non
2026-01-05 17:49:58 | IRM_Salle1.log | 6C03 | Software | MOYENNE | Oui
2025-12-08 14:27:58 | IRM_Salle1.log | 4B02 | Calibration | MOYENNE | Oui
2025-12-20 19:02:58 | Scanner_CT_01.log | 5A01 | Hardware | BASSE | Oui
2025-11-14 12:29:58 | Echo_Salle3.log | 4A01 | Calibration | MOYENNE | Non
2026-01-15 14:17:58 | IRM_Salle1.log | 3C04 | Détecteur | BASSE | Oui
2025-12-07 14:42:58 | IRM_Salle1.log | 4B02 | Calibration | HAUTE | Oui
2026-01-24 09:43:58 | Echo_Salle3.log | 7F03 | Calibration | CRITIQUE | Non
2025-12-10 18:37:58 | IRM_Salle1.log | 6C03 | Hardware | MOYENNE | Non
2026-01-29 09:27:58 | Echo_Salle3.log | 4A01 | Power | MOYENNE | Non
2026-02-04 02:02:58 | Scanner_CT_01.log | 3C04 | Hardware | BASSE | Non
2026-01-05 20:50:58 | Mammo_01.log | 4B02 | Tube RX | HAUTE | Non
2025-12-01 14:35:58 | IRM_Salle1.log | 5A01 | Calibration | BASSE | Non
2026-02-09 11:34:58 | Echo_Salle3.log | 3C04 | Détecteur | CRITIQUE | Non
2026-01-19 00:01:58 | Radio_DR_02.log | 4A01 | Network | MOYENNE | Non
2025-12-11 19:23:58 | Radio_DR_02.log | 8D01 | Hardware | BASSE | Oui
2025-12-19 12:47:58 | Radio_DR_02.log | 4B02 | Tube RX | CRITIQUE | Non
2026-02-01 22:58:58 | Mammo_01.log | 4B02 | Network | HAUTE | Oui
2025-11-18 13:31:58 | Echo_Salle3.log | 6C03 | Network | BASSE | Oui
2025-11-24 06:37:58 | Radio_DR_02.log | 7F03 | Calibration | CRITIQUE | Oui
2026-01-19 10:29:58 | Radio_DR_02.log | 4B02 | Calibration | BASSE | Oui
2025-11-23 20:52:58 | Echo_Salle3.log | 5B02 | Power | BASSE | Non
2026-02-06 23:47:58 | Radio_DR_02.log | 5A01 | Calibration | HAUTE | Oui
2025-12-17 14:59:58 | Scanner_CT_01.log | 4A01 | Calibration | BASSE | Oui
2025-12-12 03:05:58 | Mammo_01.log | 5B02 | Software | BASSE | Non
2026-01-01 18:21:58 | Radio_DR_02.log | 3C04 | Network | HAUTE | Oui
2026-02-04 23:29:58 | Echo_Salle3.log | 8D01 | Calibration | HAUTE | Oui
2025-12-05 11:58:58 | Mammo_01.log | 3C04 | Hardware | BASSE | Non
2025-11-22 03:41:58 | Radio_DR_02.log | 3C04 | Détecteur | MOYENNE | Non
2026-02-01 03:06:58 | Mammo_01.log | 6C03 | Software | CRITIQUE | Non
2026-01-26 15:44:58 | IRM_Salle1.log | 5A01 | Software | CRITIQUE | Non
2025-12-11 15:38:58 | Echo_Salle3.log | 8D01 | Power | CRITIQUE | Non
2025-12-14 19:03:58 | Radio_DR_02.log | 7F03 | Power | HAUTE | Oui
2025-11-21 22:54:58 | Radio_DR_02.log | 6C03 | Tube RX | CRITIQUE | Oui
2026-02-06 00:49:58 | Scanner_CT_01.log | 4B02 | Détecteur | BASSE | Oui
2025-11-15 14:41:58 | Mammo_01.log | 3C04 | Détecteur | MOYENNE | Non
2025-12-19 15:21:58 | Echo_Salle3.log | 5A01 | Détecteur | MOYENNE | Oui
2026-02-09 09:02:58 | Echo_Salle3.log | 4B02 | Hardware | MOYENNE | Oui
2026-02-05 08:15:58 | Scanner_CT_01.log | 6C03 | Tube RX | BASSE | Non
2026-01-16 00:28:58 | Mammo_01.log | 4B02 | Détecteur | BASSE | Non
2026-01-29 14:56:58 | Radio_DR_02.log | 3C04 | Calibration | CRITIQUE | Non
2025-12-10 07:54:58 | Mammo_01.log | 3C04 | Détecteur | CRITIQUE | Non
2026-02-09 11:13:58 | Scanner_CT_01.log | 6C03 | Software | HAUTE | Non
2025-11-24 00:22:58 | Scanner_CT_01.log | 7F03 | Software | CRITIQUE | Oui
2026-01-19 05:15:58 | IRM_Salle1.log | 4A01 | Power | BASSE | Oui
2026-02-06 14:56:58 | IRM_Salle1.log | 6C03 | Calibration | BASSE | Oui
2026-01-20 22:11:58 | Radio_DR_02.log | 7F03 | Software | HAUTE | Non
2026-02-08 15:47:58 | Mammo_01.log | 5B02 | Network | MOYENNE | Oui
2025-11-16 15:54:58 | Echo_Salle3.log | 4A01 | Power | HAUTE | Non
2025-12-31 19:18:58 | Radio_DR_02.log | 5A01 | Power | HAUTE | Non
2025-11-27 06:06:58 | Echo_Salle3.log | 5A01 | Détecteur | CRITIQUE | Oui
2025-12-13 05:40:58 | Radio_DR_02.log | 6C03 | Software | CRITIQUE | Non
2026-01-20 17:13:58 | Scanner_CT_01.log | 8D01 | Software | BASSE | Oui
2025-11-23 15:03:58 | Mammo_01.log | 4A01 | Hardware | BASSE | Oui
2025-12-02 20:17:58 | Echo_Salle3.log | 5B02 | Hardware | CRITIQUE | Oui
2026-01-02 05:17:58 | Echo_Salle3.log | 7F03 | Power | MOYENNE | Non
2026-01-22 02:59:58 | IRM_Salle1.log | 3C04 | Hardware | BASSE | Non
2025-11-27 22:39:58 | Echo_Salle3.log | 5A01 | Tube RX | MOYENNE | Non
2025-11-22 16:11:58 | Scanner_CT_01.log | 7F03 | Hardware | MOYENNE | Non
2026-02-03 03:15:58 | Radio_DR_02.log | 3C04 | Power | BASSE | Non
2026-01-27 16:14:58 | Radio_DR_02.log | 6C03 | Hardware | BASSE | Non
2025-11-19 16:24:58 | Scanner_CT_01.log | 3C04 | Software | HAUTE | Oui
2026-01-25 01:39:58 | Mammo_01.log | 5A01 | Hardware | HAUTE | Oui
2026-01-13 00:06:58 | Echo_Salle3.log | 6C03 | Power | MOYENNE | Oui
2026-01-09 09:31:58 | Radio_DR_02.log | 6C03 | Hardware | CRITIQUE | Oui
2026-02-08 23:17:58 | Mammo_01.log | 3C04 | Détecteur | MOYENNE | Oui
2025-12-25 03:01:58 | Radio_DR_02.log | 4B02 | Tube RX | HAUTE | Non
2026-02-08 00:07:58 | Echo_Salle3.log | 5A01 | Software | CRITIQUE | Oui
2025-12-18 21:37:58 | Scanner_CT_01.log | 3C04 | Détecteur | HAUTE | Oui
2025-12-28 23:01:58 | Mammo_01.log | 4A01 | Calibration | CRITIQUE | Non
2025-11-15 04:24:58 | Radio_DR_02.log | 8D01 | Détecteur | CRITIQUE | Non
2025-11-30 03:37:58 | Echo_Salle3.log | 5A01 | Détecteur | MOYENNE | Non
2025-12-17 13:18:58 | IRM_Salle1.log | 6C03 | Calibration | MOYENNE | Non
2026-01-19 09:52:58 | Scanner_CT_01.log | 5B02 | Tube RX | CRITIQUE | Oui
2026-02-03 10:53:58 | Radio_DR_02.log | 8D01 | Power | CRITIQUE | Oui
2025-12-29 20:46:58 | Radio_DR_02.log | 4A01 | Software | BASSE | Oui
2026-01-24 19:23:58 | Scanner_CT_01.log | 7F03 | Hardware | BASSE | Non
2026-01-04 20:36:58 | Scanner_CT_01.log | 4A01 | Détecteur | HAUTE | Non
2026-02-03 09:56:58 | Radio_DR_02.log | 6C03 | Calibration | CRITIQUE | Oui
2025-11-29 14:24:58 | Radio_DR_02.log | 8D01 | Power | MOYENNE | Non
2026-02-08 06:33:58 | IRM_Salle1.log | 6C03 | Software | BASSE | Non
2025-11-17 02:06:58 | Radio_DR_02.log | 5B02 | Network | CRITIQUE | Oui
2025-12-26 01:27:58 | Mammo_01.log | 5B02 | Détecteur | MOYENNE | Oui
2025-12-03 16:55:58 | IRM_Salle1.log | 5B02 | Power | CRITIQUE | Non
2026-01-07 09:15:58 | Mammo_01.log | 6C03 | Software | BASSE | Non
2026-01-06 13:17:58 | Scanner_CT_01.log | 3C04 | Software | BASSE | Non
2026-01-18 17:46:58 | Scanner_CT_01.log | 5A01 | Hardware | MOYENNE | Non
2026-01-06 14:05:58 | Scanner_CT_01.log | 4B02 | Hardware | CRITIQUE | Non
2026-02-10 22:11:24 | test_simulation_scanner.log | 2026 | Software | MOYENNE | Non
2026-02-10 22:23:55 | Install.log | DDEFA | Software | HAUTE | Non
2026-02-10 23:34:59 | Install.log | TXT | Software | BASSE | Non

## EQUIPEMENTS
Nom | Type | Fabricant | Modele | NumSerie | DateInstallation | DernieresMaintenance | Statut | Notes
Scanner_CT_01.log | Scanner CT | Siemens | SOMATOM go.Up | SN-2024-001 | 2022-03-15 | 2025-12-01 | Opérationnel | Scanner principal salle 2
IRM_Salle1.log | IRM | GE Healthcare | SIGNA Explorer | SN-2024-002 | 2021-06-20 | 2026-01-15 | Opérationnel | IRM 1.5T salle 1
Radio_DR_02.log | Radiographie Numérique (DR) | Philips | DigitalDiagnost C90 | SN-2024-003 | 2023-01-10 | 2025-11-20 | Opérationnel | Radio numérique salle urgences
Mammo_01.log | Mammographe | Hologic | Selenia Dimensions | SN-2024-004 | 2022-09-05 | 2025-10-15 | En maintenance | Mammographe numérique 3D
Echo_Salle3.log | Échographe | Samsung | RS85 Prestige | SN-2024-005 | 2024-02-28 | 2026-01-20 | Opérationnel | Échographe polyvalent salle 3